              PRESSURE    TEST    RECORD

Test Date:  6/04/07
Item Description :  Sonardyne Depth Transducer – 7970-000-02
Item S/N:  231444 - 02
Tested For: MBARI at our own Pressure Test Facilities – Autoclave Unit
Test Pressure:  4015 psi
  4015  psi / 1.5 S.F = 2677 psi    2677 / 1.487 psi/meter S.W. = 1800 meters water depth.
Gauge:  Autoclave Pressure Transducer
Gauge S/N:  95104916
Gauge Range:  0-15000 psi

Calibration Date:  8/30/06 with a PSI-Tronixs digital gage, PG-5000, 0-10ksi, S/N K0637. This unit calibrated 7/27/06, 0-10ksi +/- 0.1%FS, to ANSI / NCSL Z540.1 standards by Scientific Technologies Inc. with an AMETEK DWT, Cert.# 402503  
The test articles were subjected to the implodable items alternative procedure stipulating 10 cycles to maximum allowed pressure.  A 1 minute hold at 0 and 4015 psi on cycles 
1-9 and a 60 minute hold at 4015 psi on the 10th cycle. 
 No evidence of leaks or damage was noted.  
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