              PRESSURE    TEST    RECORD

Test Date:  4/9-10/07
Item Description :  6000m Sharkeye II, Ti case
Item S/N:  none
Tested For: Desert Star in MBARI Large Chamber
Test Pressure:  9200 psi
 9200 / 1.487 psi/meter S.W. = 6187 meters water depth.
Gauge:  Heise Dial
Gauge S/N:  H35771
Gauge Range:  0-10000 psi

Calibration Date:  8/30/06  with a PSI-Tronixs 0-10ksi digital gage, PG-5000, S/N K0637. This unit calibrated 7/27/06, 0-10ksi +/- 0.1%FS, to ANSI / NCSL Z540.1 standards by Scientific Technologies Inc. with an AMETEK DWT, Cert.# 402503  

The test articles were subjected to 3 cycles to maximum allowed pressure.  A 1 minute hold at 0 and 9200 psi on cycles 1 and  2 and a 60 minute hold at 9200 psi on the 3rd cycle.   Pressure trace is attached.
 No evidence of leaks or damage was noted.  

Test Technician:   




James T. Scholfield




MBARI




7700 Sandholdt Rd.




Moss Landing, CA.  95039




831-775-1765




scji@mbari.org

