Hopkins EnHSW system in qkFOCE mode						9/17/14

Note gauge pressures—TANK G1-0-15 PSI, Hose  G2--5-15 psi

System is currently stable with no controls and as of 9/16, fully functional.   Tank is at 13 psi,  inlet (Hopkins)  SW pressure is ~12 psi so no water enters even though valves open.  A slow flow out the hose maintains ~10 psi. As water leaves the tank (about ½ full now) the tank pressure will decrease and allow some water to enter as the pressure equalizes.  As water enters and  level rises, pressure will rise- stopping flow.

TO RUN SYSTEM;

- Water should be visible in the Tank window, black lines ( ~2” apart)show normal variation when pumps energized.

- Turn on all labeled switches. Recirc pump and light will come on immediately.  Inlet pump may or may not start.

- Open tank inlet valve, V5 one full turn,  it’s at ¼ T right now.

- Close V8  (90 deg.) and close V10.  Swap bypass for Peristaltic Pump tubing.

- Open V8 and V10.

- Turn on P.P. so get positive flow, ~ 0.5 to 1 L/m.  1 L/m is about ½ way up the glass.

- Observe functionality.  Inlet Pump should start if water Level is at or below the lower black line on the tank window. It should also stop when reaches the upper mark.

- The pressure in the hose, G2, will slowly rise to 20 – 30 psi depending on flow as the Back Press. Reg. at hose end sees enough pressure to start  regulating.

- To change to CO2 gas,  blow off tank to 0 psi, V14 on top.  Close V14, open V15 from the CO2 tanks.  Make sure CO2 reg set to 15 psi.  It will take most of a small cylinder to fill Tank to 15 psi.

- NOTE:  V6 (Hose press. Relief is not present)

- Shut down is reverse of above including shut off CO2. Don’t need to bleed tank.  Close V5 to ¼ turn. Reconnect bypass, Reset flow to 0.2 L/m.  shut off all power,