Radiometer Service Report

| Date | 4/19/18 | SR #: | 1005503068R | Sensor Type: | HyperOCR-R08W |
| --- | --- | --- | --- | --- | --- |

| S/N | 254 | Customer | MBARI |
| --- | --- | --- | --- |

Problems Found:

Failed Calibration. Dark Range was out of spec (high).

Failed Calibration again after the Spec Analog board was changed. Dark Range was 0.

Failed Calibration again after Spec Digital board was changed.

Raw Counts were unstable.

Passed Calibration after Spectrometer was changed.

Services Performed:

Did Post Calibration.

Replaced the Spec Serial Analog board. Failed Cal.

Replaced the Spec Digital board. Failed Cal.

Replaced the Spectrometer.

Passed Calibration and Reality Check.

Came in with Bioshutter 063.