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Short Term Noise:
 .881 mV

· Standard deviation of voltage output with air in the instrument’s flow tube.

· Samples taken at a one-second interval over approximately 60 seconds.

Long Term Stability:
.467 mV/Hour

· Deviation of voltage output over a 12-hour period with air in the flow tube.

· One sample/minute for 12 hours. Ambient temperature is constant.

Full Scale Voltage:
5.456 Volts

Temperature Stability:
59.0 mV @
24oC


42.0 mV @
2oC

· Voltage output of the instrument with water in the flow tubes. Instrument is placed in a water bath and the ambient temperature is recorded. 

Instrument Response:
Pure water
0.068 Volts


copro.1
0.5 mg/l (nominal)    2.996 Volts

· Voltage output of the instrument with sample in the flow tube. The output is average voltage over 30-second interval, one sample per second.   

1copro. = coproporphyrin tetramethyl ester concentration: 0.5 mg/l of copro. is approximately equal to 50 ug/l of chlorophyll in a Thaalssiosira weissflogii phytoplankton culture.
Template saved in F:\Controlled Documentation\Characterization\WETStar\wschar.doc


[image: image1.jpg]