{"text_format":"1","workbook":"xlsx","n_sheets":3,"sheets":["Volumes","equations","Sheet3"],"not_rendered":{"charts":0,"images":0,"macros":1,"external_links":0,"defined_names":0,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"Volumes","index":0,"n_rows":32,"n_cols":12,"rows_kept":32,"cut":false,"hidden":false,"merged":[]}
["Sample bottle",null,null,null,null,null,"TItrant bottle"]
["ml","temp (deg C)","mass","volume","volume at 20 deg",null,"ml","temp (deg C)","mass","volume","volume at 20 deg"]
[29.915400000000002,17.821999999999999,{"$f":"A3*(1-((3.4848*(101.325-(0.003796*30*EXP(-5315.56/(B3+273.15))*(1.7526*10^8)))/(273.15+B3))*10^-3)/8)/(1-((3.4848*(101.325-(0.003796*30*EXP(-5315.56/(B3+273.15))*(1.7526*10^8)))/(273.15+B3))*10^-3)/((999.84847+B3*6.337563*10^-2-B3^2*8.523829*10^-3+B3^3*6.943248*10^-5-3.821216*10^-7*B3^4)/1000))","v":29.947180329874044},{"$f":"C3/((999.84847+B3*6.337563*10^-2-B3^2*8.523829*10^-3+B3^3*6.943248*10^-5-3.821216*10^-7*B3^4)/1000)","v":29.988412419440543},{"$f":"D3*(1+((1+32.5*10^-7)^3-1)*(20-B3))","v":29.989049240442139},null,0.99890000000000001,18.417999999999999,{"$f":"G3*(1-((3.4848*(101.325-(0.003796*30*EXP(-5315.56/(H3+273.15))*(1.7526*10^8)))/(273.15+H3))*10^-3)/8)/(1-((3.4848*(101.325-(0.003796*30*EXP(-5315.56/(H3+273.15))*(1.7526*10^8)))/(273.15+H3))*10^-3)/((999.84847+H3*6.337563*10^-2-H3^2*8.523829*10^-3+H3^3*6.943248*10^-5-3.821216*10^-7*H3^4)/1000))","v":0.99995904164084926},{"$f":"I3/((999.84847+H3*6.337563*10^-2-H3^2*8.523829*10^-3+H3^3*6.943248*10^-5-3.821216*10^-7*H3^4)/1000)","v":1.0014471170704011},{"$f":"J3*(1+((1+32.5*10^-7)^3-1)*(20-H3))","v":1.0014625639416608},"2017-11-13"]
[29.954799999999999,17.866,{"$f":"A4*(1-((3.4848*(101.325-(0.003796*30*EXP(-5315.56/(B4+273.15))*(1.7526*10^8)))/(273.15+B4))*10^-3)/8)/(1-((3.4848*(101.325-(0.003796*30*EXP(-5315.56/(B4+273.15))*(1.7526*10^8)))/(273.15+B4))*10^-3)/((999.84847+B4*6.337563*10^-2-B4^2*8.523829*10^-3+B4^3*6.943248*10^-5-3.821216*10^-7*B4^4)/1000))","v":29.98661746005169},{"$f":"C4/((999.84847+B4*6.337563*10^-2-B4^2*8.523829*10^-3+B4^3*6.943248*10^-5-3.821216*10^-7*B4^4)/1000)","v":30.028146210021827},{"$f":"D4*(1+((1+32.5*10^-7)^3-1)*(20-B4))","v":30.028770992676488},null,0.996,18.373999999999999,{"$f":"G4*(1-((3.4848*(101.325-(0.003796*30*EXP(-5315.56/(H4+273.15))*(1.7526*10^8)))/(273.15+H4))*10^-3)/8)/(1-((3.4848*(101.325-(0.003796*30*EXP(-5315.56/(H4+273.15))*(1.7526*10^8)))/(273.15+H4))*10^-3)/((999.84847+H4*6.337563*10^-2-H4^2*8.523829*10^-3+H4^3*6.943248*10^-5-3.821216*10^-7*H4^4)/1000))","v":0.99705612347022754},{"$f":"I4/((999.84847+H4*6.337563*10^-2-H4^2*8.523829*10^-3+H4^3*6.943248*10^-5-3.821216*10^-7*H4^4)/1000)","v":0.99853155146181949},{"$f":"J4*(1+((1+32.5*10^-7)^3-1)*(20-H4))","v":0.99854738173321889}]
[29.9039,17.922000000000001,{"$f":"A5*(1-((3.4848*(101.325-(0.003796*30*EXP(-5315.56/(B5+273.15))*(1.7526*10^8)))/(273.15+B5))*10^-3)/8)/(1-((3.4848*(101.325-(0.003796*30*EXP(-5315.56/(B5+273.15))*(1.7526*10^8)))/(273.15+B5))*10^-3)/((999.84847+B5*6.337563*10^-2-B5^2*8.523829*10^-3+B5^3*6.943248*10^-5-3.821216*10^-7*B5^4)/1000))","v":29.935657392725506},{"$f":"C5/((999.84847+B5*6.337563*10^-2-B5^2*8.523829*10^-3+B5^3*6.943248*10^-5-3.821216*10^-7*B5^4)/1000)","v":29.977424440419096},29.978031800000888,null,1.0064,18.367999999999999,{"$f":"G5*(1-((3.4848*(101.325-(0.003796*30*EXP(-5315.56/(H5+273.15))*(1.7526*10^8)))/(273.15+H5))*10^-3)/8)/(1-((3.4848*(101.325-(0.003796*30*EXP(-5315.56/(H5+273.15))*(1.7526*10^8)))/(273.15+H5))*10^-3)/((999.84847+H5*6.337563*10^-2-H5^2*8.523829*10^-3+H5^3*6.943248*10^-5-3.821216*10^-7*H5^4)/1000))","v":1.007467172824283},{"$f":"I5/((999.84847+H5*6.337563*10^-2-H5^2*8.523829*10^-3+H5^3*6.943248*10^-5-3.821216*10^-7*H5^4)/1000)","v":1.0089568611663997},1.0089729157401519]
[29.87,18.224,29.901689106482973,29.945090573993092,29.945609104866694,null,0.99370000000000003,18.312000000000001,0.99475390464387003,0.99621425484776827,0.99623065059526072]
[29.805199999999999,18.204999999999998,29.836822384537932,29.880023228981543,29.880546168437633,null,0.99050000000000005,18.292999999999999,0.99155057799596236,0.99300266890733058,0.99301919575271203]
[29.869900000000001,18.312000000000001,29.901579607851396,29.945476673923068,29.945969518179911,null,0.9738,18.187000000000001,0.97483323399856625,0.97624140666269832,0.9762586634940682]
[29.8688,18.167999999999999,29.900493813766975,29.943579411356446,29.944114265310162,null,0.98740000000000006,18.199000000000002,0.98844762168328026,0.98987768558097033,0.98989506764215129]
[29.9069,18.254999999999999,29.938624940117204,29.982254584319403,29.982764696561201,null,0.98550000000000004,18.337,0.98654511983981574,0.98799808129344335,0.98801410099339704]
[29.9147,18.405999999999999,29.9464170815735,29.990913157638484,29.991379262930167,null,0.9919,18.236999999999998,0.99295226139937498,0.99439593054794484,0.99441302352374605]
[29.907399999999999,18.349,29.939115427757649,29.983277069523051,29.983759719398471,null,0.99080000000000001,18.187000000000001,0.99185127156066888,0.99328402723495723,0.99330158532544943]
[29.899000000000001,21.253,29.930400007532601,29.992369112149721,29.992002701683536,null,0.99309999999999998,21.507999999999999,{"$f":"G13*(1-((3.4848*(101.325-(0.003796*30*EXP(-5315.56/(H13+273.15))*(1.7526*10^8)))/(273.15+H13))*10^-3)/8)/(1-((3.4848*(101.325-(0.003796*30*EXP(-5315.56/(H13+273.15))*(1.7526*10^8)))/(273.15+H13))*10^-3)/((999.84847+H13*6.337563*10^-2-H13^2*8.523829*10^-3+H13^3*6.943248*10^-5-3.821216*10^-7*H13^4)/1000))","v":0.99414207302251756},{"$f":"I13/((999.84847+H13*6.337563*10^-2-H13^2*8.523829*10^-3+H13^3*6.943248*10^-5-3.821216*10^-7*H13^4)/1000)","v":0.99625660226013768},{"$f":"J13*(1+((1+32.5*10^-7)^3-1)*(20-H13))","v":0.99624195425170869},"2017-11-16"]
[29.924600000000002,21.338000000000001,29.956018016293072,30.018602573086572,30.018210964133974,null,0.99580000000000002,21.344000000000001,0.99684547552505076,0.99892942377790084,0.99891633376418909]
[29.916599999999999,21.393000000000001,29.948003878009818,30.010936705424665,30.01052910306036,null,0.99609999999999999,21.308,0.99714591560210242,0.9992225516658223,0.99920980853920582]
[29.915099999999999,21.422999999999998,29.946499174203495,30.009628324319909,30.009211961880954,null,0.99519999999999997,21.478000000000002,0.99624438061641907,0.99835671875988663,0.99834233189363364]
[29.930599999999998,21.544,29.962002822448333,30.025972496547478,30.025520484088474,null,0.99419999999999997,21.574999999999999,0.99524299523701198,0.99737476462000407,0.99735944863399795]
[29.920300000000001,21.542999999999999,29.951692120034288,30.015633087827833,30.015181523672172,null,0.99670000000000003,21.611000000000001,0.99774549299803716,0.99989066704537422,0.99987496146165111]
[29.9221,21.504000000000001,29.953498074045608,30.017182108629243,30.016741935240237,null,0.99970000000000003,21.611000000000001,1.0007486398616814,1.0029002707386985,1.0028845178822241]
[29.912700000000001,21.49,29.944089669596522,30.007660228986602,30.007224291285826,null,0.99519999999999997,21.611000000000001,0.99624391956621505,0.99838586519871209,0.99837018325136462]
[29.933700000000002,21.611000000000001,29.965099090754837,30.029524691618459,30.029053008833777,null,0.99239999999999995,21.611000000000001,0.99344098249348056,0.99557690175160962,0.99556126392549671]
[29.933900000000001,21.52,29.965308787422899,30.029124908035094,30.0286798749576,null,0.99029999999999996,21.805,0.99133811129908411,0.99351280965503963,0.99349532501465565]
[]
[]
[]
[null,null,null,null,{"$f":"AVERAGE(E3:E22)","v":29.991117530882029},null,null,null,null,null,{"$f":"AVERAGE(K3:K22)","v":0.99601856386799725}]
[null,null,null,null,{"$f":"_xlfn.STDEV.P(E3:E22)","v":3.6898325383655693E-2},null,null,null,null,null,{"$f":"_xlfn.STDEV.P(K3:K22)","v":6.3828617797878429E-3}]
[null,null,null,null,{"$f":"E27/E26","v":1.2303084520162101E-3},null,null,null,null,null,{"$f":"K27/K26","v":6.4083763208190227E-3}]
[]
[null,null,null,null,{"$f":"AVERAGE(E3,E5:E6,E8:E12)","v":29.970084700961205},null,null,null,null,null,{"$f":"AVERAGE(K3:K7,K11:K22)","v":0.99801196736648989}]
[null,null,null,null,{"$f":"_xlfn.STDEV.P(E3,E5:E6,E8:E12)","v":1.9618605011116536E-2},null,null,null,null,null,{"$f":"_xlfn.STDEV.P(K3:K7,K11:K22)","v":3.8748359937853163E-3}]
[null,null,null,null,{"$f":"E31/E30","v":6.5460625843634422E-4},null,null,null,null,null,{"$f":"K31/K30","v":3.8825546391092515E-3}]

{"sheet":"equations","index":1,"n_rows":16,"n_cols":6,"rows_kept":16,"cut":false,"hidden":false,"merged":[]}
["Computaiton of air density",null,null,"Computation of mass from weight"]
["p=air pressure",101.325,null,"m=w(1-p(air)/p(weight))/(1-p(air)/p(sample))"]
["U=humidity",30,null,"w=100",100,100]
["t=temperature",20,{"$f":"F3*(1-((3.4848*(101.325-(0.003796*30*EXP(-5315.56/(B4+273.15))*(1.7526*10^8)))/(273.15+B4))*10^-3)/8)/(1-((3.4848*(101.325-(0.003796*30*EXP(-5315.56/(B4+273.15))*(1.7526*10^8)))/(273.15+B4))*10^-3)/((999.84847+B4*6.337563*10^-2-B4^2*8.523829*10^-3+B4^3*6.943248*10^-5-3.821216*10^-7*B4^4)/1000))","v":100.10545974387765},"p(air)=.0012013",{"$f":"((3.4848*(B2-(0.003796*B3*B5))/(273.15+B4))*10^-3)","v":1.2013290002401886E-3}]
["es=vapor pressure",{"$f":"EXP(-5315.56/(B4+273.15))*(1.7526*10^8)","v":2.3378250253394524},null,"p(weight)=8",8,8]
["p(air)=.0012013",{"$f":"((3.4848*(B2-(0.003796*B3*EXP(-5315.56/(B4+273.15))*(1.7526*10^8)))/(273.15+B4))*10^-3)","v":1.2013290002401886E-3},null,"p(sample)=1",1,1]
[]
[null,null,null,"m",{"$f":"E3*(1-B6/8)/(1-B6/B13)","v":100.10545974387765}]
["Computation of water pressure at temperature"]
["p(H2O)=(999.84847+temp*6.337563*10^-2-temp^2*8.523829*10^-3+temp^3*6.943248*10^-5-3.821216*10^-7*temp^4)/1000"]
[5,{"$f":"(999.84847+A11*6.337563*10^-2-A11^2*8.523829*10^-3+A11^3*6.943248*10^-5-3.821216*10^-7*A11^4)/1000","v":0.99996069265900001}]
[10,{"$f":"(999.84847+A12*6.337563*10^-2-A12^2*8.523829*10^-3+A12^3*6.943248*10^-5-3.821216*10^-7*A12^4)/1000","v":0.99969545466400012},null,"av",{"$f":"(1+32.5*10^-7)^3-1","v":9.7500316875720472E-6}]
[20,{"$f":"(999.84847+A13*6.337563*10^-2-A13^2*8.523829*10^-3+A13^3*6.943248*10^-5-3.821216*10^-7*A13^4)/1000","v":0.9982007713840001},null,null,{"$f":"3*32.5*10^-7","v":9.7499999999999998E-6}]
[23,{"$f":"(999.84847+A14*6.337563*10^-2-A14^2*8.523829*10^-3+A14^3*6.943248*10^-5-3.821216*10^-7*A14^4)/1000","v":0.99753485564249444},{"$f":"30*(1+E12*(23-20))","v":30.000877502851882},{"$f":"0.997541-B14","v":6.1443575055664112E-6}]
[30,{"$f":"(999.84847+A15*6.337563*10^-2-A15^2*8.523829*10^-3+A15^3*6.943248*10^-5-3.821216*10^-7*A15^4)/1000","v":0.995643451264},null,{"$f":"D14/E12","v":0.63018846527425787}]
[40,{"$f":"(999.84847+A16*6.337563*10^-2-A16^2*8.523829*10^-3+A16^3*6.943248*10^-5-3.821216*10^-7*A16^4)/1000","v":0.99221081622399998}]

{"sheet":"Sheet3","index":2,"n_rows":0,"n_cols":0,"rows_kept":0,"cut":false,"hidden":false,"merged":[]}