## Slide 1: Photosynthetic Response Incubation and Manipulation System (PRIMaS)
Zbigniew Kolber,
Ocean Sciences, University of California, Santa Cruz, California 95064
zkolber@ucsc.edu

## Slide 2
PRIMaS System

## Slide 3
Temperature Control
0-35C By TE
Cooling/Heating
Light Control
0-300Wt/m2
LED Cluster
Blue, Green, White
Media:
Pumping in
Pumping out
0-3 L/24H
pH Control
CO2/Air mixture
pH: 6.0 – 8.3
Dissolved Oxygen
Nitrogen/Air mixture
Stirring
Magnetically
coupled
Design
Incubation vessel
custom design
D100mm, H300mm
Borosilicate glass
(SIMAX)

## Slide 4
Gas Control System
+12V power supply
NI Compact field point controller
CO2 Flow controller
+12V
+12V
N2 Flow controller
Air Flow controller
+12V
4-20mA
Air 30-40PSI
N2 30-45PSI
CO2
Gas dispersion tube, 4-8 um
Common manifold
0-10cc/min
0-50cc/min
0-100cc/min
Filter

## Slide 5
pH Control/Monitoring Sysytem
+12V power supply
NI Compact field point controller
pH (ORP) transmitter
+12V
CO2 Flow controller
CO2
Common manifold for remaining gases
4-20 mA
pH electrode

## Slide 6
Media Flow Control
2-channel peristaltic pump
Extraction tube (bigger)
Delivery tube (smaller)
Sterile air filter
FRRF with flow-thru cell
Fresh
media
Spend
media
Sterile air filter

## Slide 7
Vessel Design
Ultrasoft heat transfer material
(used for power electronics backplane)
Jar, D100mm, H300mm, Borosilicate glass (SIMAX).
Welded piece of glass pipe and round bottom.
Tight diameter tolerance
Aluminum jacket
(as opposed to double-wall vessel)

## Slide 8: Experimental Setup
Agilent Technologies GC (Model 6850, Series II)
Reduced gas analyzer (RGA) (Peak Laboratories, Mnt View, CA, USA)
PRIMAS Incubation System
FRR Fluorometer

## Slide 9: Daily patterns of photosynthetic activity in Crocosphaera
Light ON
Light OFF
Phot. Yield
Fluorescence
Efficiency of light utilization
Electron transport time
Mixed open ocean populations

## Slide 10
Temporal patterns of photosynthetic activity
and nitrogen fixation in Crocosphaera

## Slide 11: Effects of pH Changes

## Slide 13
PRIMaS
  Control of temperature, light, nutrients, pCO2, pH, and DO
  Measurements of pH, DO, optical density
Fast Repetition Rate Fluorometry
(FRRF)
  Fluorescence signal
  Fv/Fm
  Functional absorption cross section
  Kinetics of electron transport
  Size of the PQ pool

## Slide 14
Relationship Between Fluorescence and Photosynthesis

## Slide 15
Excitation trace
Fluorescence emission trace
Ratio of integrated emission to integrated excitation signal (fluorescence yield)

## Slide 16
Fast Repetition Rate (FRR) Method

## Slide 18
Rapid Light Curves (RLC)
R rate of electron transport
E irradiance
sPSII functional absorption cross section of PSII
Fm maximum level of fluorescence yield
Fs fluorescence yield measured under ambient irradiance.
E
R
Ek
a

## Slide 19
Light
ETR

## Slide 20
Instantaneous P vs I curves
It works… sometimes

## Slide 21
20 mM NH4CL
20 mM DBMIB
Effects of Inhibitors on Nitrogen Fixation in Crocosphaera

## Slide 22
FD
PSII
2e-
PQ
PQH2
Cyt b6f
PC
2e-
2H+
NAD(P)H NAD(P)+
PSI
e-
e-
e-
FNR
Calvin Cycle
FD
FD
F1
6H+
ADP
ATP
2H2O O2+4H+
FQR
Fo
ATP, CO2
ADP
H+
H+
H+
Photosynthetic Electron Flow
Lumen
GLY
PYR
TCA
Cycle
GLU
G3P
ADP
ATP, CO2
OPP
Cycle

## Slide 23
FD
FD?
O2+4H+ 2H2O
PSII
2e-
PQ
PQH2
HupL
HupS
HupC
Cyt b6f
PC
2e-
2H+
NAD(P)H NAD(P)+
H+
e-
PSI
e-
e-
2e-
2H+
FNR?
e-
Nitrogenase
Net H2
e-
NAD(P)+ NAD(P)H
FNR
8e-
2 NH3
GLY
PYR
Calvin Cycle
“TCA”
Cycle
GLU
FPOR?
e-
G3P
H2
FD
FD
Ndh
e-
F1
6H+
ADP
ATP
Cyt Oxidase
2H2O O2+4H+
FQR
Fo
ATP, CO2
ADP
ADP
ATP, CO2
Lumen
OPP
Cycle
e-
H+
H+
Photosynthetic/Respiratory Electron Flow